## Slide 2: عدم قطعیت در آزمایشگاه بالینی
برآورد عدم قطعیت بر اساس راهنمای ISO/TS 20914:2019
حسن بیات
دانش آموخته علوم آزمایشگاهی

## Slide 3: What’s the Problem?
“Although all medical laboratories seeking ISO 15189:2012 accreditation know that MU estimate is a specific requirement, few know what to do with the calculated MU.”
Braga F, Panteghini M. The utility of measurement uncertainty in medical laboratories. Clin Chem Lab Med 2020
“We wish to add that “few are truly aware of how to calculate MU in medical laboratories”.
Mario Plebani, Andrea Padoan and Laura Sciacovelli Measurement uncertainty: light in the shadows https://doi.org/10.1515/cclm-2020-0134

## Slide 4: Gaussian Distribution
-∞
+∞
(µ)
SD (σ)
Normal (Gaussian) distribution
Notes: aaa

## Slide 5
-∞
+∞
Mean
SD
1.5
Example
Body Mass Index (BMI)
Mean = 23 Kg/m2
SD = 1.5 Kg/m2
Frequency of Over-weight (BMI>25)?
Frequency of Obese (BMI>30)?
25
Overweight
30
Obese
9.1%
0.002%
Probability of BMI>25?
Probability of BMI>30?

## Slide 6
Note: Irrespective of the shape of the curve, area under curve is a function of the number of SDs

## Slide 7
Example
Repeated assay of Glucose:
True Value = 100 mg/dL
SD = 3 mg/dL
The middle 68% of results width?
The middle 95% of results width?
100
103
106
109
91
94
97
Probability of falling between 97 and 103?
Probability of falling between 94 and 106?

## Slide 8
Q: With 95% probability, where is a certain bird?
A: In the range -2SD to +2SD from the seeds place.
+2SD
-2SD

## Slide 9
Q: With 95% probability, where is the seeds place?

## Slide 10
Q: With 95% probability, where is the seeds place?

## Slide 11
Q: With 95% probability, where is the seeds place?

## Slide 12
Q: With 95% probability, where is the seeds place?

## Slide 13
Q: With 95% probability, where is the seeds place?

## Slide 14
Q: With 95% probability, where is the seeds place?
+2SD
-2SD
A: In the range -2SD to +2SD from the bird.

## Slide 15
Example
Repeated assay of Glucose:
True Value = 100 mg/dL
SD = 3 mg/dL
100
103
106
109
91
94
97

## Slide 16
Q: With 95% probability, what is the True Value?
+2SD
-2SD
y
Result
TV
TV

## Slide 17
With 95% probability?
y - 2SD
y ± 2SD
y
y + 2SD
95% probability
True Value:
With 68% probability?
y ± 1SD
68% probability
y - 1SD
y + 1SD
With 99.7% probability?
y ± 3SD
99.7% probability
y - 1SD
y + 1SD
y ± 1SD

## Slide 18: Definition
Measurement uncertainty, MU:
Parameter characterizing the dispersion of the quantity values being attributed to a measurand, based on the information used.
VIM – 3rd Ed. GCGM 2012
It describes the interval within which the value of the measurand is assumed to lie with a stated level of confidence

## Slide 19: Why MU?
To ensure that measurement results are useful and safe in medical practice:
Meaningful comparison with medical decision limits
Meaningful comparison with previous results
ISO 15189:2012,
5.5.1.4: “…shall determine measurement uncertainty for each measurement procedure in the examination phase used”
Screening
Diagnosis
Classification
Monitoring

## Slide 20: Data sources
Type A: Evaluation of MU from the frequency distribution
Type B: Evaluation of MU based on experience or other information
Standard Deviation
&
Gaussian (Normal) distribution

## Slide 21: Approaches to estimating uncertainty
Bottom-up:
- Determining different sources of uncertainty
- Establish a mathematical
- Estimating partial uncertainties
- Combining under the established model
Top-down:
Experimental evaluation of uncertainty in the final outcome
Bottom-up
Top-down

## Slide 22: Example
Driving time: About 3 hrs
Uncertainty

## Slide 23: Example
Bias Correction?

## Slide 24
Uncertainty of arrival time
Bus status
Weather
Traffic
Driver
Accident
Passengers
Clock Adjusting
Fishbone Diagram

## Slide 25
Bottom-up
Hey! Listen to my experience.
I’ve taken this trip many times…
Top-down

## Slide 26
Sources of uncertainty in Medical Laboratory
Question: Bottom-up or Top-down?
Bias Correction
Control Material

## Slide 27
Evaluation of measurement data — Guide to the expression of uncertainty in measurement (GUM) JCGM 100:2008
1.2: “This document is primarily concerned with the expression of uncertainty in the measurement of a well-defined physical quantity that can be defined by an essentially unique value.”
1.4: “…(GUM) provides general rules for evaluating and expressing uncertainty in measurement rather than detailed, technology-specific instructions.”

## Slide 28
Bottom-up:
Usually employed by reference laboratories to obtain ISO 17025 and 15195 certificate
The application in medical laboratories is too complicated and has encountered many practical problems and objections.
Top-down:
Simpler and represents a good alternative to the previous approach
Estimates MU by using IQC, commercial calibrator information, and bias correction
The search for any potential source of analytical uncertainty according to the “bottom-up” approach is seen as a complicated and useless challenge
Measurement uncertainty: light in the shadows. Plebani et al. CCLM 2020; aop. https://doi.org/10.1515/cclm-2020-0134
The utility of measurement uncertainty in medical laboratories. Braga et al. CCLM 2020; aop. https://doi.org/10.1515/cclm-2019-1336

## Slide 30
ISO/TS 20914:2019(E)
In medical laboratory:‍
The measurands are typically present in complex biological fluid and tissue matrices
Characterization of commercial devices and kits is limited to the gathering of empirical performance data using QC sample
Medical laboratory measurement procedures are well-suited to utilizing IQC and other available data to estimate MU without the need for measurement models and complex statistics
“…not intended … an exhaustive accounting of all possible uncertainties, but … major foreseeable sources…”
“…to provide practical estimates of MU that help ensure that patient results are fit for medical use.”

## Slide 31
Common sources of MU in medical laboratory:
Sample inhomogeneity
Reconstitution procedures for lyophilized materials, e.g. calibrator and reagents
Uncertainty of calibrator values, re-calibrations
Instrumentation, e.g. electro-mechanical fluctuations, maintenance, parts replacement
Reagent and calibrator instability
Reagent and calibrator lot-to-lot variability
Fluctuations in laboratory environment
Operator bias introduced by reading analogue instrument indications
Variable manual dexterity for manual and semi-automated methods
Measurement bias vs. an accepted calibration hierarchy scheme
Measurement formulae, e.g. approximations, assumptions, inexact values for constants, rounding of digits
More than one of the same measuring system for the same measurand
More than one measurement procedure for the same measurand that may have different analytical performance characteristics

## Slide 32
The most significant uncertainty contributions
1. Long-term imprecision data (SD or CV) obtained for IQC
Inhomogeneity
Recalibrations
Reagent/Calibrator lot changes
Instrumentation (e.g. electro-mechanical, maintenance, part replacement
Environment
Operators
Formula (e.g. approximation, assumptions, rounding)
More than one system
More than one measurement procedures
Uncertainty of within-lab reproducibility, uRw

## Slide 33
An SD from such long-term precision data is termed a standard uncertainty (u)
If a period sufficient to include all routine changes, the u is designated uRw, or the standard uncertainty under within-laboratory conditions
Example:
Na IQC results

## Slide 34
The most significant uncertainty contributions
1. Long-term imprecision data (SD or CV) obtained for IQC
2. Uncertainty of end-user calibrator values, ucal
3. Occasionally a medically significant bias may be corrected, ubias
uRw ucal ubias
u(y) = √(u2Rw + u2cal + u2bias)
Unfortunately, the information regarding uncertainty associated with commercial calibrators is currently very poor when consulting calibrator package inserts.
In the majority of cases, long-term imprecision is the largest contributor

## Slide 35: 1. Long-term imprecision, uRw
In the majority of cases, the largest contributor to the uncertainty
IQC materials close to important medical decision limits
IQC data produced when the measuring system is meeting analytical performance criteria
More than one level of IQC
Dissolving a lyophilized IQC material introduces a spurious contribution to the estimation of MU
Preferably by a third-party (i.e. different from that used to check the alignment of the measuring system)
Closely resembles authentic clinical samples (ideally a commutable material)
If IQC material does not undergo preparation step, a typical uncertainty for this step must be estimated

## Slide 36: NOTE: Changes during IQC data collection
Calibrator lot change/Recalibratoin
Maintenance/Part replacement
Reagent lot change
IQC lot change
Chart: IQC
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
 | 125 | 122 | 122 | 122 | 121 | 122 | 122 | 121 | 125 | 124 | 121 | 123 | 123 | 123 | 123 | 124 | 123 | 121 | 125 | 124 | 132 | 133 | 135 | 135 | 132 | 134 | 132 | 135 | 134 | 131 | 134 | 135 | 132 | 135 | 134 | 134 | 131 | 134 | 131 | 134
Single set;
Larger SD
Avg. separate sets
Smaller SD
Avg. separate sets;
Smaller SD

## Slide 37: NOTE: Changes during IQC data collection
Calibrator lot change/Recalibratoin
Maintenance/Part replacement
Reagent lot change
IQC lot change
YES. Usually affects IQC as patient samples. Treat IQC data as a single set
YES/NO. Both IQC and patient samples affected: treat IQC data as a single set; Only IQC affected: treat IQC data as separate sets
NO. Treat IQC data as separate sets
Chart: IQC
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
 | 125 | 122 | 122 | 122 | 121 | 122 | 122 | 121 | 125 | 124 | 121 | 123 | 123 | 123 | 123 | 124 | 123 | 121 | 125 | 124 | 132 | 133 | 135 | 135 | 132 | 134 | 132 | 135 | 134 | 131 | 134 | 135 | 132 | 135 | 134 | 134 | 131 | 134 | 131 | 134
Are IQC results and patient sample results affected the same way?

## Slide 38: 2. Uncertainty of end-user calibrator values, ucal
combination of all uncertainties from the highest available reference (calibrator or measurement procedure) down to the assigned value of the calibrator
Manufacturer, is responsible for determining ucal
If laboratory prepares the calibrator, the laboratory is responsible for determining ucal
For some calibrators a comprehensive ucal is technically unavailable; e.g. some international conventional biological reference materials

## Slide 39: 3. Uncertainty of bias, ubias
IVD manufacturers are responsible for ensuring minimal, medically acceptable measurement bias
Therefore estimation of bias and ubias by the laboratory is rarely required
If ongoing EQA indicates the introduction of a medically significant:
Responsibility of the IVD manufacturer, to take immediate corrective action
If the manufacturer is unable to correct bias, the laboratory may, if local regulations permit, can estimate bias by using reference materials
NOTE! EQA target value can itself be biased, depending on how the value is assigned to the material

## Slide 40
Bias correction is always imperfect!
    Uncertainty of Reference; uref
    Uncertainty of Mean; SEM or SDmean
BIAS = Lab Mean - Reference Value

## Slide 41: Caveats about correcting bias
Only category I/II EQA are usable to evaluate bias
I: Commutable, Value assigned with RMP or CRM, Replicate samples
II: Commutable, Value assigned with RMP or CRM
Correction factor not permitted in some countries because:
- May alter the status of the system, removing any responsibility from the manufacturer
- Depriving the system and the results of the CE marketing certification
May labs continue to use the correction factor even when the bias has been corrected in the reagent production stage
“Therefore, we are not supporting the individual use of bias correction factors in daily practice”
The utility of measurement uncertainty in medical laboratories. Braga et al. CCLM 2020; aop. https://doi.org/10.1515/cclm-2019-1336
III: Commutable, Replicate samples
IV: Commutable
V: Non-commutable, Replicate samples
VI: Non-commutable

## Slide 42
If local regulations permit, laboratory can manage bias by:
Applying a correction factor to the results
Re-assigning a calibrator value
Bias correction is always imperfect because estimates of the magnitude of bias have uncertainty (ubias)
Bias is not the case when:
- Monitor changes in an individual over time
- In-house determined reference intervals

## Slide 43: MU is all about TRACEABILITY
Example of a calibration hierarchy for a measurand with full metrological traceability to the SI
Uncertainty of Reference Material; uref
Uncertainty of End-user Calibrator; ucal
Uncertainty of Patient result; u(y)
uRw

## Slide 44: MU is all about TRACEABILITY
Example of a calibration hierarchy for a measurand with full metrological traceability to the SI

## Slide 45: Maximum allowable measurement uncertainty
The magnitude of MU should be suitable for a result to be used in a medical decision and ideally as small as technically possible
If a measurement procedure exceeds the upper allowable uncertainty limit, it may be possible to reduce the uncertainty by identifying and modifying a technical step that makes a relatively large contribution to the total uncertainty
If the uncertainty cannot be sufficiently reduced, a decision must be made as to whether the measurement procedure is to be changed or if the quality goal should be reconsidered
Allowable CV; I(%)
Optimum: 0.25 x CVI
Desirable: 0.5 x CVI
Minimum: 0.75 x CVI
CVI
CVG

## Slide 46
Standard measurement uncertainty; u
Measurement uncertainty expressed as a standard deviation, SD
ur: Repeatability uncertainty
uRw: Within-lab Reproducibility uncertainty
Coverage factor, k
Number larger than one by which a standard uncertainty value (u) is multiplied to obtain an expanded uncertainty, U
Expanded uncertainty; U U = k * u
Measurement uncertainty obtained from multiplying standard uncertainty by coverage factor
Coverage interval [(x – U) to (x + U)]
Interval containing the set of true quantity values of a measurand with a stated probability, based on the information available
Concepts & Definitions

## Slide 47
Metrological Traceability
Property of a measurement result whereby the result can be related to a reference through a documented unbroken chain of calibrations, each contributing to the MU
Standard Error, SE
Quantitative measure of the variation or dispersion of sample means or sample averages
EXAMPLE: A reference material for serum creatinine was measured 20 times under repeatability conditions:
- Mean value = 22,0 N = 20 SD = 0,63
- SDmean
- Coverage Interval = [21.7 – 22.3]; 95% Confidence
= 0.63/√20 = 0.15 u = 0.15
U = 0.3, k = 2

## Slide 48: Basic calculations
Variance and Standard Deviation
Combining independently estimated uncertainties
NOTE! Variance can be added, subtracted, multiplied or divided, but SD cannot

## Slide 49
Rounding
Performed only on the final calculated estimates
Number of decimal digits:
- Mean: Plus one additional decimal digit
- SD: Plus two additional decimal digits
- CV: One decimal digits
“The number of significant digits in the results shall not imply a smaller uncertainty than warranted by the uncertainty.”

## Slide 50
Example:
Sodium measuring system reports Na values in whole numbers
136, 134, 133, 135, 136, 137, 135, 138, 136, 133
 = 135.3
SD = 1.64
CV = 1.2%

## Slide 51: Re-estimation of measurement uncertainty
Apart from updating interim estimates, there is no need to re-estimate MU if IQC and EQA performance remain within specifications.
Re-estimation of MU should be undertaken if there has been a significant change in a measuring system or a new measurement procedure is introduced.

## Slide 53
Flowchart Diagram for calculating MU

## Slide 54: Example - Uncertainty of Sodium

## Slide 56: Example - Uncertainty Intraoperative PTH monitoring Calculation of ur(y) under Repeatability conditions
Not applicable to the uncertainty for long-term precision conditions for other medical uses of PTH values. Useful as initial guidance.

## Slide 57: MU across changes affecting only IQC results
Chart: IQC
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
 | 125 | 122 | 122 | 122 | 121 | 122 | 122 | 121 | 125 | 124 | 121 | 123 | 123 | 123 | 123 | 124 | 123 | 121 | 125 | 124 | 132 | 133 | 135 | 135 | 132 | 134 | 132 | 135 | 134 | 131 | 134 | 135 | 132 | 135 | 134 | 134 | 131 | 134 | 131 | 134
Calibrator lot change/Recalibratoin
Maintenance
Reagent lot change
IQC lot change
Usually affects IQC as patient samples. Treat IQC data as a single set
- Both IQC and patient samples affected: treat IQC data as a single set
- Only IQC affected: treat IQC data as separate sets
Treat IQC data as separate sets

## Slide 58: MU across changes affecting only IQC results
Under circumstances where there are very large differences in the number of IQC data points (n) between different measuring conditions:
Chart: IQC
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40
 | 125 | 122 | 122 | 122 | 121 | 122 | 122 | 121 | 125 | 124 | 121 | 123 | 123 | 123 | 123 | 124 | 123 | 121 | 125 | 124 | 132 | 133 | 135 | 135 | 132 | 134 | 132 | 135 | 134 | 131 | 134 | 135 | 132 | 135 | 134 | 134 | 131 | 134 | 131 | 134
SD1 = u1
SD2 = u2

## Slide 59: Example for Multiple IQC Lots

## Slide 60
Under circumstances where there are very large differences in the number of IQC data points (n) between different measuring conditions:

## Slide 61: Example – Uncertainty of PTH across 3 IQC lots

## Slide 65
average = (1 + 2 + 3)/3 | 2.13667 | 17.87333 | 61.57

## Slide 66: Laboratories using multiple measuring systems for the same measurand
Average u that can be applied to the results irrespective of the measuring system used.
uRw must be estimated separately for each measuring system, and then a pooled average can be calculated:
  Average within system Uncertainty, uRw(A,B,C)
Medically significant systematic bias between measuring systems should be eliminated, and the uncertainty of the mean residual bias should be calculate:
Between systems Uncertainty, u(A,B,C)
Total Uncertainty:

## Slide 67
Chart: C
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
 | 121 | 122 | 120 | 118 | 121 | 123 | 120 | 119 | 122 | 119 | 120 | 120 | 121 | 118 | 120 | 118 | 122 | 122 | 123 | 123
Chart: B
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
 | 128 | 127 | 127 | 122 | 124 | 129 | 128 | 124 | 129 | 127 | 125 | 123 | 129 | 126 | 122 | 123 | 128 | 128 | 124 | 123
 | 128 | 127 | 127 | 122 | 124 | 129 | 128 | 124 | 129 | 127 | 125 | 123 | 129 | 126 | 122 | 123 | 128 | 128 | 124 | 123
Chart: A
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
 | 122 | 124 | 124 | 121 | 121 | 125 | 125 | 122 | 124 | 124 | 123 | 122 | 124 | 125 | 124 | 123 | 121 | 122 | 122 | 125
Chart: A,B,C
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
 | 122 | 124 | 124 | 121 | 121 | 125 | 125 | 122 | 124 | 124 | 123 | 122 | 124 | 125 | 124 | 123 | 121 | 122 | 122 | 125
 | 128 | 127 | 127 | 122 | 124 | 129 | 128 | 124 | 129 | 127 | 125 | 123 | 129 | 126 | 122 | 123 | 128 | 128 | 124 | 123
 | 128 | 127 | 127 | 122 | 124 | 129 | 128 | 124 | 129 | 127 | 125 | 123 | 129 | 126 | 122 | 123 | 128 | 128 | 124 | 123
 | 121 | 122 | 120 | 118 | 121 | 123 | 120 | 119 | 122 | 119 | 120 | 120 | 121 | 118 | 120 | 118 | 122 | 122 | 123 | 123
 A, SD A
 C, SD C
 B, SD B

## Slide 68
 | Measuring System |  | 
 | A | B | C
n | 280 | 190 | 400
 | 5.15 | 4.93 | 5.28
SD | 0.160 | 0.190 | 0.200
uRW | 0.160 | 0.190 | 0.200
u(A,B,C) = SD(A, B, C) = 0.176918
u2(A,B,C) = 0.0313
 avg = 5.12

## Slide 69: Example – Uncertainty of WBC counts

## Slide 73: Standard uncertainty (u) vs. Relative uncertainty (urel)
A. Standard uncertainty method
Example - Albumin in Serum/Plasma

## Slide 76
Minimum

## Slide 77
B. Relative uncertainty method

## Slide 79
Minimum

## Slide 80
Relative uncertainty method
Standard uncertainty method
Relative uncertainty method gives a better estimate of overall uncertainty
Standard uncertainty method underestimates overall uncertainty at levels higher than calibrator concentration

## Slide 82
Multipoint Calibration
For practical purposes, the calibrator that is most applicable to each level of IQC is used for ucal

## Slide 84: Bias, Bias Correction, Uncertainty of bias
The role of IVD manufacturers is to define a calibration hierarchy to assign traceable values to their calibrators
If the IVD measuring system is CE-marked and correct alignment to higher-order references is expected:
Medical laboratories should rely on the IVD manufacturers who must ensure traceability
Therefore, regular estimation of bias by the end-user laboratory is not required
The utility of measurement uncertainty in medical laboratories. Braga et al. CCLM 2020; aop. https://doi.org/10.1515/cclm-2019-1336

## Slide 85
Example
Lab
n = 10
 = 108.9
Glucose CRM
Assigned vale = 108 mg/dL
uref = 0.6 mg/dL
Bias = 108.9 - 108= 0.9 mg/dL
Significant Bias?
AV – Uref AV + Uref
108
106.8
109.2
Uref= 1.2 mg/dL
CI = 106.8 to 109.2 mg/dL
lab
108.9
Not statistically significant; correction not necessary.

## Slide 86
Example
Lab
n = 10
 = 109.6
SD = 1.3
SDmean (SE) = 1.3/√10 = 0.41
2SE = 0.82
CImean = 108.8 to 110.4 mg/dL
Glucoscante CRM
Assigned vale = 108 mg/dL
uref = 0.6 mg/dL
Bias = 109.6 - 108= 1.6 mg/dL
Mean – 2SE mean + 2SE
Significant Bias?
AV – Uref AV + Uref
108
106.8
109.2
Uref= 1.2 mg/dL
CI = 106.8 to 109.2 mg/dL
lab
109.6
110.4
108.8

## Slide 87
Example
Lab
n = 10
 = 109.6
SD = 1.3
SDmean (SE) = 1.3/√10 = 0.41
Glucoscante CRM
Assigned vale = 108 mg/dL
uref = 0.6 mg/dL
Bias = 109.6 - 108= 1.6 mg/dL
Significant Bias?
ubias = √(uref2 + SDmean2)
If Bias > 2ubias → Statistically Significant Bias
ubias = √(0.62 + 0.412) = 0.73 2ubias = 1.46
1.6 > 1.46
Statistically Significant
Medically Significant Bias?

## Slide 88
Bias > 2ubias
Statistically Significant Bias
NOTE!
The final assessment of significance of any correction for the bias should be based on its impact on medical interpretation of measurement results.
BIAS

## Slide 89
Example
Lab
n = 10
 = 3.72
SD = 0.11
SDmean = 0.11/√10 = 0.034785
CRM (Albumin)
Assigned vale = 4.15 mg/dL
uref = 0.040
ubias = √(0.042 + 0.0347852 ) = 0.053009
2ubias 0.11
Bias = 4.15 – 3.72 = 0.43 > 2ubias

## Slide 90
Component (Analyte) | Albumin
Measurand | Amount of albumin concentration in serum/plasma
Measurand unit | mg/dL
Bias (mg/dL) | Assessed by EQA using commutable material with target value set by accredited reference laboratory: 0.43; Significant bias; Correction required
Bias uncertainty, uBias | 0.053009
Calibrator uncertainty, ucal | 0.058
IQC mean, mg/dL | 4.22
Laboratory uncertainty, uRw | 0.077918
u albumin | √(0.0530092 + 0.0582 + 0.0779182) = 0.110658 ≈ 0.11
U albumin | 0.221316 ≈ 0.22

## Slide 91
شاد باشید

## Slide 92: Combined uncertainty of calculated values
Examples:
Anion Gap
Cr Clearance
Calculated LDL
Non-HDL Cholesterol
ACR

## Slide 93
1- Measured value is calculated by adding/subtracting, e.g. LDL:
Combine Standard uncertainties
2- Measured value is calculated by Multiplying/Dividing; e.g. Cr Clearance:
Combine Relative uncertainties

## Slide 94: 1- Adding/Subtracting
Example 1 – Uncertainty of Plasma Anion Gap

## Slide 95: 1- Adding/Subtracting
Example 2 – Uncertainty of Osmolality
Na = 130, u(Na) = 0,98
Urea = 6.5, u(urea) = 0,19
Glucose = 5.2, u(glu) = 0,090

## Slide 96
2- Multiplying/Dividing
Example 1– Uncertainty of Spot Urine Calcium to Creatinine ratio
Ca = 6.40, u(Ca) = 0,040 (IQC mean: 2.71)
Cr = 2.30, u(Cr) = 0,150 (IQC mean: 6.17)
Ca/Cr = 2.78
%urel(Ca) = (0.04/2.71) x 100 = 1.4760 % = 1.5 %
%urel(Cr) = (0.15/6.17) x 100 = 2.4311 % = 2.4 %
%urel(Ca/Cr) = √(1.4762 + 2.43112) = 2.8394 % = 2.8 %
%Urel(Ca/Cr) = 2 x 2.8394 % = 5.7 %
Ca/Cr = 2.78 ± 5.7% = [2.62 – 2.93]

## Slide 97
2- Multiplying/Dividing
Example 2– uncertainty of Creatinine Clearance
u(P-Cr) = 1.438 (IQC mean = 70)
%urel(Cr) = (1.438/70) x 100 = 2.05 %
u(Ur-Cr) = 138.0 (IQC mean = 6060)
%urel(Ur-Cr) = (138/6060) x 100 = 2.28%
Notes: To convert u to %u-rel, u is divided by IQC mean

## Slide 98: Rectangular & Triangular distributions

## Slide 99
2- Multiplying/Dividing
Example 1– uncertainty of Creatinine Clearance
u(V) = 100/√3 = 57.7
%urel(V) = (57.7/2421) x 100 = 2.40 %
Volume Error = ± 100 mL
u(t) = 30/√3 = 17.32
%urel(t) = (17.32/1441) x 100 = 1.20%
Time Error = ± 30 min

## Slide 100
2- Multiplying/Dividing
Example 1: Creatinine Clearance
%urel(V) = 2.40 %
%urel(t) = 1.20%
Clearance = 33.4
%urel(Cr) = 2.05 %
%urel(Ur-Cr) = 2.28%
%urel(Clearance) = 2.05 % = 4.0744 % = 4.1%
%Urel(Clearance) = 8.1488 = 8.1%
33.4 x (8.1488/100) = 2.7217 = 2.72
Patient result = 33.4 ± 2.72 = [30.7 – 36.1]

## Slide 101
INR = (PT/MNCT)ISI
SD of IQC
SDmean of Normal Specimens
2- Multiplying/Dividing
Example 2– Uncertainty of INR

## Slide 103
ISI
urel MNCT
Urel (Rw)(PT)

## Slide 104: Uncertainty of Calculated results: Direct estimation using IQC data
Better reflect the uncertainty
Include, for example, small uncertainties introduced by automatic rounding of values by a laboratory IT system

## Slide 105
Example 1: Anion Gap = (Na + K) – (Cl + HCO3)
SD (AG)
 (AG)

## Slide 106
Example Patient:
[Na+] = 143 mmol/l
[K+] = 4,0 mmol/l
[Cl-] = 104 mmol/l
[HCO3-] = 22 mmol/l
AG = (143 + 4) – (104 + 22) = 21 mmol/l
Patient AG is between 18 and 24; 95% confidence
%U (AG) = (3/25.85) x 100 = 11.6%

## Slide 107: Example 2: MU of eGFR
For method standardized against NIST reference method/material

## Slide 108
Calculate eGFR for phantom 40 years old female and male using IQC results
Female
QC1
Male
QC1
Female
QC2
Male
QC2

## Slide 109
0.93 mg/dL
0.02 mg/dL
6.03 mg/dL
0.10 mg/dL

## Slide 110
0.93 mg/dL
0.02 mg/dL
6.03 mg/dL
0.10 mg/dL
Example- Female; eGFR = 80
U = (80 x 5.2)/100 = 4.16
Patient eGFR is between 76 and 84; 95% confidence

## Slide 111
0.93 mg/dL
0.02 mg/dL
6.03 mg/dL
0.10 mg/dL
Example- Male; eGFR = 11
U = (11 x 3.7)/100 = 0.41
Patient eGFR is between 10.6 and 11.4; 95% confidence

## Slide 112
NOTE!
The best-fit function to derive CKD-EPI eGFR reveals very significant scatter
eGFR values could be up to approximately ±30 % of the measured GFR value
Therefore, uncertainty from laboratory performance is only a small portion of Total Diagnostic Uncertainty!

## Slide 113
“Doctors continually have to make decisions on the basis of imperfect data and limited knowledge, which leads to diagnostic uncertainty, coupled with the uncertainty that arises from unpredictable patient responses to treatment and from health care outcomes that are far from binary.”
Han PK, Djulbegovic B. Tolerating uncertainty about conceptual models of uncertainty in health care. J Eval Clin Pract 2019;25:183–5
William Osler:
“Medicine is a science of uncertainty and an art of probability”
NOTE!
The best-fit function to derive CKD-EPI eGFR reveals very significant scatter
eGFR values could be up to approximately ±30 % of the measured GFR value
Therefore, uncertainty from laboratory performance is only a small portion of Total Diagnostic Uncertainty!

## Slide 114
SUMMARY:
uRw ucal ubias
u = √(uRw2 + ucal2 + ubias2)
u = √(uRw2 + ucal2)
u = uRw
u < CVI → Acceptable
U = 2u
Result ± U

## Slide 115
According to ISO 15189 [63], MU should be made available by the laboratory on request. Careful interpretation is needed, as MU is already taken into account also in other test characteristics such as sensitivity, specificity, predictive value, etc. Physicians might correct decision levels knowing the MU of a test, taking MU into account again. Analytical variation as well as other sources of uncertainty, including biological and preanalytical variation, are already reflected in reference interval limits and are (or should be) also taken into account in decision limits in clinical guidelines, as “grey zones”, and borderline areas. Lower MU results in higher predictive values and lower clinical uncertainty. This might even lead to a different clinical application of the test, with HbA1c assays as a typical example. The improved reliability of the results made expert organizations revise the guidelines and recommend HbA1c for diagnosis. Such an improved clinical uncertainty is also seen with high-sensitive troponin assays.
Careful interpretation is needed
…are already reflected in RIs and decision limits as “grey zones” and borderline areas.

## Slide 116
شاد باشید

## Slide 117: MU calculation according to different purposes
*For tests used for both monitoring and diagnosis the specific clinical question should be considered
Measurement uncertainty: light in the shadows. Plebani et al. CCLM 2020; aop. https://doi.org/10.1515/cclm-2020-0134

## Slide 118: Example - Uncertainty of BCR-ABL1 gene transcript measurement

## Slide 120
Example - Uncertainty of HIV viral load

## Slide 121
Note! Non-asymmetric confidence interval

## Slide 122: Uncertainty of Qualitative results based on numerical results
Delineate ‘uncertainty zones’ around the negative/positive cut-offs, e.g. negative, probably negative, probably positive, positive
At signal values other than those near the decision limits is unnecessary

## Slide 123
Example - Uncertainty of HBs Ag

## Slide 124: Bias, Bias Correction, Uncertainty of bias
Example of a calibration hierarchy for a measurand with full metrological traceability to the SI

## Slide 125: Bias, Bias Correction, Uncertainty of bias
Example of a calibration hierarchy for a measurand with full metrological traceability to the SI