## Slide 1
عیار سیگما
در مرحله سنجش در آزمایشگاه بالینی
حسن بیات
درس خوانده علوم آزمایشگاهی

## Slide 2: Background
Six Sigma management technology:
Developed by Motorola and General Electric in the 1980s
To systematically improve processes and eliminate defects
Sigma Metric is a core concept in Six Sigma Methodology

## Slide 3: SM – Industrial view
Upper SL
Lower SL
TV - S
Centered performance
Specification, S (allowable deviation)
Specification Limit
TV
Target Value, TV
TV + S

## Slide 4: SM – Industrial view
Target Value, TV
Lower SL
TV - S
TV
TV + S
Upper SL
Defect
Defect
Perfect
Centered performance
Specification, S (allowable deviation)
Specification Limit

## Slide 5: SM – Industrial view
USL
LSL
TV
Capability:
Cp=
(USL – LSL)/6SD

## Slide 6: SM – Industrial view
USL
LSL
TV
Capability:
Off-center
Cp=
(USL – LSL)/6SD

## Slide 7: SM – Industrial view
USL
LSL
TV
Capability:
Off-center
Cp=
(USL – LSL)/6SD

## Slide 8
5 meter
5 meter
3 meter
7 meter
7 meters away…; I’m really safe!
Just 3 meters!
2 m

## Slide 9: SM – Industrial view
USL
LSL
TV
Capability:
Cpk = min
Cp=
D/3SD
D/3SD
D/3SD

## Slide 10: SM – Industrial view
USL
LSL
TV
Capability:
Cpk = min
Cp=
D/3SD
D/3SD

## Slide 11: SM – Industrial view
USL
LSL
TV
Capability:
Cpk = min
Cp=
D/3SD
D/3SD

## Slide 12: SM – Industrial view
USL
LSL
TV
Capability:
Cpk = min
Cp=
D/3SD
D/3SD
If 3 SDs feet between mean and SL:
Defect = 0.03%
= 300 DPM
Cpk = 1

## Slide 13
SIGMA METRIC
Derived from Cpk
Combines “spread” and “centeredness”
Measure of correct outputs
USL
LSL
TV

## Slide 14
SIGMA METRIC
Derived from Cpk
Combines “spread” and “centeredness”
Measure of correct outputs
USL
LSL
TV

## Slide 15
SIGMA METRIC
Derived from Cpk
Combines “spread” and “centeredness”
Measure of correct outputs
USL
LSL
TV
Number of Sigmas (SDs) between “Mean” and nearest SL
How many SDs?

## Slide 16
USL
LSL
TV
SM =
How many SDs?
SIGMA METRIC
Derived from Cpk
Combines “spread” and “centeredness”
Measure of correct outputs
Number of Sigmas (SDs) between “Mean” and nearest SL

## Slide 17
عیار سیگما و شمارش خطا
رویکرد شمارشی
رویکرد آماری (توزیع-بنیان)

## Slide 18
روش شمارش

## Slide 19
روش آماری (توزیع-بنیان)

## Slide 20
  روش آماری
  نمونه‌گیری از محصول
  برآورد عدم صحت و عدم دقت
  محاسبه سطح زیر منحنی بیرون از مرزهای مجاز

## Slide 21
Excellent
World Class
Good
Marginal
Poor
Unacceptable
SM reflects defect rate

## Slide 22: SM – Medical Laboratory
In medical laboratory:
Allowable Total Error (ATE) equals industrial Specification
Bias equals industrial Shift (Off-centeredness)
SM =

## Slide 23: SM – Medical Laboratory
SM =
SM =
Introduced by James O. Westgard
In medical laboratory:
Allowable Total Error (ATE) equals industrial Specification
Bias equals industrial Shift (Off-centeredness)

## Slide 24: SM – Medical Laboratory
ATE
Bias
SD

## Slide 25: SM – Medical Laboratory
ATE
Bias
SD

## Slide 26: SM – Medical Laboratory
ATE
Bias
SD
Estimated:
Certified Reference Materials - JCTLM, IFCC, NIST, WHO,…
EQA/PT materials
QC & Calibration Verification materials
Spiked samples
Conventional TVs
Patient samples tested in one or more central labs
Patient samples sent to a traceable lab traceable
Previously collected QC data
CLSI EP09
CLSI EP15

## Slide 27: SM – Medical Laboratory
ATE
Bias
SD
CLSI EP05
CLSI EP15
Estimated from Verification/Validation studies
Recalculated from long-term IQC data

## Slide 28
ATE = 1.5 th/uL
Bias = 0 th/uL
SD = 0.3 th/uL
11.5
8.5
0.3 DPM
0.3 DPM
5 σ
5 σ
TV
Mean
Example: WBC assay
SM =

## Slide 29
ATE = 1.5 th/uL
Bias = 0.6 th/uL
SD = 0.3 th/uL
11.5
8.5
0.000001 DPM
1350 DPM
Example: WBC assay
SM =

## Slide 30
The number of SDs from the target to the ULs or LLs gives the SM value [2]
+ATE
-ATE
TV
How many SD?
How many SD?
SM =
Cp=

## Slide 31
The number of SDs from the target to the ULs or LLs gives the SM value [2]
+ATE
-ATE
TV
How many SD?
How many SD?
SM =
Cp=

## Slide 32
The number of SDs from the target to the ULs or LLs gives the SM value [2]
+ATE
-ATE
TV
How many SD?
How many SD?
SM =
Cp=

## Slide 34
High level (211.3 IU/L)
Bias = 17.7 IU/L
SD = 2.8 IU/L
Xc = 200 IU/L
ATE = 30 IU/L

## Slide 35
High level (211.3 IU/L)
Bias = 17.7 IU/L
SD = 2.8 IU/L
Xc = 200 IU/L
ATE = 30 IU/L
+200
-30
TV
200
BIAS

## Slide 36
High level (211.3 IU/L)
Bias = 17.7 IU/L
SD = 2.8 IU/L
Xc = 200 IU/L
ATE = 30 IU/L
+200
-30
TV
200
5.4 DPM
BIAS

## Slide 37
+ATE
-ATE
TV
How many SD?
How many SD?
SM =
Cp=
The number of SDs from the target to the ULs or LLs gives the SM value [2]

## Slide 38
Warren Brussee:
Subtract the average you just calculated from the nearest specification limit. Divide the result by the sigma σ you just calculated to get the number of sigma that “fit.” This number indicates how well the process run when it is in control.
Brussee W. Statistics for Six Sigma made easy. New York: McGraw-Hill; 2004
+ATE
-ATE
TV

## Slide 39
Warren Brussee:
Subtract the average you just calculated from the nearest specification limit. Divide the result by the sigma σ you just calculated to get the number of sigma that “fit.” This number indicates how well the process run when it is in control.
Brussee W. Statistics for Six Sigma made easy. New York: McGraw-Hill; 2004
+ATE
-ATE
TV
How many SDs?

## Slide 40
Poster #132
AACC 2022

## Slide 41: SM – Applications in medical laboratory
Quality Indicator
Goal Setting
Method Evaluation
IQC Planning
EQAS
E.g., ISO 15189: % yield, % defects, DPMO, Six Sigma scale
E.g., IFCC WG-A1C: 2 sigma for routine testing, 4 sigma for clinical trials
Individualized QC
E.g., The Netherlands SKML
Benchmarking new methods

## Slide 42
Six Sigma has brought no genuinely new and palatable food to the table of laboratory medicine
Incorrect adaptation of Six Sigma from industry to laboratory medicine and the utilisation of erroneous equations

## Slide 43
Total Analytical Error model (TAE)
Six Sigma model
Lab med vs industry
Utilization of SM in medical laboratory
Sigma metric is powerful tool for patient safety

## Slide 44
خصوصیات عیار سیگما

## Slide 45
TV
0.3 DPM
Sigma Metric = 5
SD
MEAN
DR = 0.6 DPM
0.3 DPM
Sigma Metric
1) رابطه نرخ خطا با عیار سیگما

## Slide 46
TV
6210 DPM
MEAN
5 Sigma
رابطه بین عیار سیگما و میزان خطا خطی نیست
SM = 2.5
B = 2.5 SD
1) رابطه نرخ خطا با عیار سیگما
SM: 5 → 2.5
DR: 0.6 → 6210
x 10500

## Slide 47
Sigma | DPM
1 | 317000
2 | 45500
3 | 27000
4 | 63
5 | 0.6
6 | 0.002
7 | 0.0000013

## Slide 48
TV
32
DPM
SM: 4 → 3
X̄
3 σ
1
σ
1350
DPM
DR: 32 → 1350
TV
0.0000013
DPM
SM: 7 → 6
X̄
6 σ
1
σ
0.0002
DPM
DR: 0.0000013→ 0.0002

## Slide 49
Normal distribution is not uniform, therefore moving the curve, dramatically changes the AUC within TL.
Argument: In contrast to normal distribution, bias can be treated as a linear parameter only in uniform [rectangular] distribution.

## Slide 50
TV
MEAN
SM = 5
2) عملیات حساب روی عیار سیگما

## Slide 51
TV
SM2 = SM1 - Bias
SM2 = 5 – 2 = 3
MEAN
SM = 3
BIAS
2 SD

## Slide 52
TV
SM2 = SM1 - Shift
SM2 = 3.5 – 1.5 = 2
MEAN
SM = 3.5
1.5 SD
2
1.5 SD

## Slide 53
TV
MEAN
SM = 6
SD

## Slide 54
TV
SD2 = 3 x SD1
SM2 = SM1 ÷ 3
SM2 = 6 ÷ 3 = 2
MEAN
- عدم صحت یا جابجایی میانگین، به واحد SD را می‌توان با عیار سیگما جمع و تفریق کرد
- می‌توان عیار سیگما را در فاکتور تغییر SD (یا CV) ضرب یا تقسیم کرد
SM = 2
SD2 = 3 x SD1

## Slide 55
Bias = 0
کالیبر مجدد
SM = 4.7
3.35 + 1.35
TEa = 8
SD = 1.7
Bias = -2.3 = 1.35 SD
SM = 3.35
DR = 404 DPM
3.35 SD
4.70 SD
1.35
SD
DR = 2.6 DPM
1.4 برابر افزایش عیار سیگما
155 مرتبه کاهش در نرخ خطا
مثال
Argument: Bias cannot be treated as a linear parameter.

## Slide 56
SD = 2.1
بهبود دقت
SM = 3.62
2.38 x 1.52
2.38
TEa = 8
SD = 3.2
Bias = -0.4 = 0.12 SD
SM = 2.38
DR = 17313 DPM
F = 3.2/2.1 = 1.52
3.62
DR = 296 DPM
1.52 برابر افزایش عیار سیگما
58 مرتبه کاهش در نرخ خطا
مثال

## Slide 57
TV
Defect Rate
Sigma Metric = 0
MEAN
50%
وقتی سیگما صفر است، نرخ خطا نزدیک به 50% است، و بازده اندکی بیش از 50% است (یعنی تقریبا 50% محصول سالم داریم)
TEa
Bias = TEa
3) سیگمای صفر
SM =
SM =

## Slide 58
TV
Defect Rate
MEAN
~50%
Bias = TEa
Argument: When Bias=TEa, Sigma becomes zero whereas yield is still about 50%; therefore this Sigma equation is wrong
~50%
3) سیگمای صفر

## Slide 59
TV
Defect Rate
MEAN
SM
بازده منفی معنی ندارد، اما سیگمای منفی معنی دارد!
سیگمای منفی یعنی نرخ خطا بیش از 50% است و بازده کمتر از 50%
> 50%
TEa
Bias > TEa
4) سیگمای منفی
NEGATIVE Sigma Metric
SM =
SM =
SM = -1.3

## Slide 60
TV
Defect Rate
MEAN
> 50%
Bias > TEa
Argument: Negative Sigma is meaningless because number of perfects cannot be less than zero.
<50%
4) سیگمای منفی

## Slide 61
Example 1:
TV=100 mg/dL
TEa = 8 mg/dL = 8%
Mean = 100 mg/dL
Bias = 0 mg/dL = 0%
SD = 2 mg/dL; CV = 2%
SM = (TEa - B)/SD
SM = (8 - 0)/2 = 4
SM = (TEa% - B%)/CV
SM = (8% - 0)/2% = 4
5) کدام فرمول: (TEa - B)/SD یا (TEa% - B%)/CV
TEa% = (TEa/TV) x 100
B% = (B/TV) x 100
CV = (SD/Mean) x 100

## Slide 62
Example 2:
TV=100
TEa = 8 = 8%
Mean = 103
Bias = 3 = 3%
SD = 2; CV = (2 / 103) x 100 = 1.94%
SM = (TEa - B)/SD
SM = (8 - 3)/2 = 2.50
SM = (TEa% - B%)/CV
SM = (8% - 3%)/1.96 = 2.58
TEa% = (TE/TV) x 100
B% = (B/TV) x 100
CV = (SD/Mean) x 100
5) کدام فرمول: (TEa - B)/SD یا (TEa% - B%)/CV

## Slide 63
Example 3:
TV=100
TEa = 8 = 8%
Mean = 97
Bias = -3 = -3%
SD = 2; CV = 2.06%
SM = (TEa - B)/SD
SM = (8 - 3)/2 = 2.50
SM = (TEa% - B%)/CV
SM = (8% - 3%)/2.06 = 2.43
TEa% = (TE/TV) x 100
B% = (B/TV) x 100
CV = (SD/Mean) x 100
5) کدام فرمول: (TEa - B)/SD یا (TEa% - B%)/CV

## Slide 64
TV | TEa | X̄ | Bias | SD | CV | SM = (TEa-B)/SD | SM = (TEa%-B%)/CV
100 | 8 | 100 | 0 | 2 | 2 | 4 | 4
100 | 8 | 103 | +3 | 2 | 1.94 | 2.50 | 2.58
100 | 8 | 97 | -3 | 2 | 2.06 | 2.50 | 2.43
=
اگر عدم صحت داشته باشیم، نتایج این دو فرمول کمی متفاوت است؛ در حضور عدم صحت، سیگمای حاصل از فرمول SD درست است
TEa% = TEa x 100/TV
Bias% = Bias x 100/TV
CV= SD x 100/X̄
≠
≠
5) کدام فرمول: (TEa - B)/SD یا (TEa% - B%)/CV

## Slide 65: SM formulas: Unit-based vs Percent-based
Without bias, SM from both formulas is the same
If bias, percent-based SM deviates from unit-based SM
Positive bias → Overestimate
Negative bias → Underestimate
Correction for bias needed
SM =
SM =

## Slide 66
Example 2:
TV=100
TEa = 8 = 8%
Mean = 103
Bias = 3 = 3%
SD = 2; CV = 1.94%
SM = (TEa - B)/SD
SM = (8 - 3)/2 = 2.50
SM = (TEa% - B%)/CV
SM = (8% - 3%)/1.96 = 2.58
5) کدام فرمول: (TEa - B)/SD یا (TEa% - B%)/CV

## Slide 67
Example 3:
TV=100
TEa = 8 = 8%
Mean = 97
Bias = -3 = -3%
SD = 2; CV = 2.06%
SM = (TEa - B)/SD
SM = (8 - 3)/2 = 2.50
SM = (TEa% - B%)/CV
SM = (8% - 3%)/2.06 = 2.43
5) کدام فرمول: (TEa - B)/SD یا (TEa% - B%)/CV

## Slide 68
>>>32
Example
TEa = 6
Bias = 2
SD = 1
SM = 4
P(Z>4) = 32 DPM
P(Z>2.5) = 6210
6) کوتاه مدت در مقابل بلند مدت

## Slide 69
سیگمای محاسبه شده گزارشی از گذشته است!
ما نیاز داریم که آینده را پیش‌بینی/تضمین کنیم
نظر به این که:
عیار سیگما در ارزیابی کوتاه مدت تعیین می‌شود
عوامل مختلف در کوتاه مدت کاملا بررسی نخواهد شد
برنامه‌های پایش کیفیت نمی‌توانند تغییرات جزئی را شناسایی کنند
برای بلند مدت، نرخ خطای بالاتری در نظر گرفته می‌شود
پاسخ
6) کوتاه مدت در مقابل بلند مدت

## Slide 70
مفروضات روش شش سیگما:
در بلند مدت، کم و بیش جابجایی کالیبراسیون در هر دو جهت اتفاق می‌افتد
این جابجایی‌ها کوچکتر از حساسیت پایش کیفیت هستند و شناسایی نمی‌شود
حتا در صورت شناسایی، اصلاح آن‌ها عملا ممکن نیست
این جابجایی‌ها خودبخود اصلاح می‌شود
حداکثر جابجایی، یعنی بدترین وضعیت، برابر 1.5 SD در نظر گرفته می‌شود
بر اساس قرارداد بنا شده در موتورولا، برای محاسبه نرخ خطای آینده (بلند مدت)، 1.5 SD از SM محاسبه شده از دادهای کوتاه مدت کم می‌شود
در روش شش سیگما برای پیش‌بینی نرخ خطا در بلند مدت، پس از محاسبه z از داده‌های کوتاه مدت، 1.5 واحد از آن کم می‌شود و سپس میزان خطا محاسبه می‌شود
6) کوتاه مدت در مقابل بلند مدت

## Slide 71
Example
TEa = 6
Bias = 2
SD = 1
SM = 4
P(Z>4) = 32 DPM
32 DPM
Target
-TEa
+TEa
Mean
Bias
2SD
Shift
1.5 SD
Worst Scenario: Shift = 1.5 SD
STABLE
SM
STABLE
Defect Rate
SM = 4 - 1.5 = 2.5
DR = 6210 DPM
6210 DPM

## Slide 72
عیار سیگمای محاسبه شده با دو نرخ خطا مربوط است:
کوتاه مدت: عملکرد دیده شده
بلند مدت: عملکرد قابل تضمین
Sigma | DPM Short term | DPM Long term
1 | 317000 | 698000
2 | 45500 | 309000
3 | 27000 | 67000
4 | 63 | 6000
5 | 0.6 | 233
6 | 0.002 | 3.4
نرخ خطای بلند مدت در آزمایشگاه، بستگی دارد به توانایی برنامه پایش کیفیت
Argument: Bias shouldn’t be incorporated in Sigma equation because the presence of 1.5 SD bias is accepted in advance.

## Slide 73
شاد باشید...